## Slide 1
Gamma spectrometric measurements of uranium isotopic composition:
-Accuracy and precision
H. Ramebäck, P. Lagerkvist, S. Holmgren, S. Jonsson, B. Sandström, A. Tovedal, A. Vesterlund, T. Vidmar#, J. Kastlander
FOI CBRN Defence and Security, Sweden
# SCK-CEN, Belgium
Presentation@GAMMASPEC 2016

## Slide 2: The problem

## Slide 3: Agenda
Gamma spectrometric measurements of uranium
Theory
Experimental
True coincidence summation
Results
Conclusions

## Slide 4: Gamma spectrometry of uranium isotopics
In this presention the use of the ’high energy’ region is considered, i.e. 120-1001 keV!
In general, GS is a fast and ’non-destructive’ measurement method
However, somewhat high uncertainties: -Uf235=10-20%, k=2, in this work
Dedicated commercial softwares available, but in this work an in house algorithm was used

## Slide 5: Theory: Isotopics using HRGS
Count rate at Eg,j from measurement!
Isotope ratio(s) as measurand(s)!
I and l from DDEP!
-Fit Ri och c1-c5 by means of the least square method Uncertainties using the Jackknife method!
Ai
The efficiency:
Isotope ratio:
The relative efficiency:
Empirical response fcn:

## Slide 6: Theory
The isotope ratios Ri: -R234=n234/n238 -R235=n234/n238
The abundance of 235U: -f235=R235/SRi (Observe: R238=1) where SRi=1+R234+R235
Measurands!

## Slide 7: Experimental
Samples: Three LEU samples and IRMM-184 (NU CRM): -One LEU and IRMM-184 as acid solutions (measured in the standard geometry: 60 mL sample container) -Two LEU as UO2 pellets
Detector: -p-type HPGe (~75×75 mm) -Semi-empirical calibration: possible to calculate efficiencies for non-std geometries, as well as different correction factors
Measurements directly on detector encapTrue coincidence summations!?
TCS correction factors calculated using EFFTRAN and VGSL

## Slide 8: True coincidence summation (TCS)
Gamma photons emitted within the time resolution of the detector system: Count losses… Depends on the geometry (solid angle): -Detector size -Sample size -Detector-to-sample distance Also detector type… The effect cancels out when the sample is far away from the detector (distant geos), but that will reduce the sensitivity, resulting in much longer counting times!
(In this work: reducing the TCS on 258 keV from 234mPa ’enough’ would require a distance which reduces the efficiency by about a factor of 15…)
Classical example:

## Slide 9: TCS: Uranium
No significant TCS for 235U in this work
Some peaks from 238U (234mPa!) largely affected
258 keV the most affected peak!!! This peak is the most important one for uranium isotopics (in particular for low enrichments) when using the ’high energy region’, and a systematic effect for this peak will result in a sytematic effect for the enrichment (with about the same factor)

## Slide 10: The correction factor
Adding the correction factor for TCS, kTCS.i,Ej, to the measurement model:

## Slide 11: Results: LEU dissolved in the ’std-geo’
Calculated response for the particular geometry
After correction
!!!

## Slide 12: Results: LEU as an uranium pellet
After correction
Calculated response for the particular geometry
!!!

## Slide 13: Results: 235U

## Slide 14: Resultat: 238U

## Slide 15: Resultat: 234U

## Slide 16: Conclusions
Without TCS correction the enrichments of uranium samples were overestimated (maximum about 40% in this work) using the detector and geometries as in this work
However, corrections resulted in excellent agreement with reference values (MS and a ref mtrl), i.e. no significant deviations
Maximum bias was <4%, and Uf235=10% (k=2), for the LEU materials

## Slide 17
FOI CBRN Defence and Security, Umeå
?

## Slide 18: Complementary information

## Slide 19: 235U abundance (FRAM evaluations included):