## Slide 1
Quality assurance of a characterised Ge-detector used for various geometries
(mathematical efficiency calibration)
Asser Nyander Poulsen
Special Adviser, Ph.D.
apo@sis.dk
Strålebeskyttelse i Sundhedsstyrelsen (SIS)
Knapholm 7
2730 Herlev, Denmark
www.sis.dk

## Slide 2
Setup:
Characterized Ge detector
Electrical cooling
Shielded lab + 10 cm Pb.
No background peaks
except 511 keV

## Slide 4
Can You prove it?
It is so…
Notes: QA objectives:
Minimized manual data handling (typing) – less QC work.
All entries and selections needed for reproduction must be recorded in a secure way
(time stamped and locked)
Either specific values or references to QA documentation (e.g. standard geometry)
Customized report output integrating system validation info. Optimized for QC.

## Slide 5: Human factor

## Slide 6: To reproduce the analysis, I need:

## Slide 7
Transfer
Manuel selection
Manual entry
Non-editable software
Editable spreadsheet
.ECC
File
.CNF
.NLB
.CAL
.GEO
Genie2K
Geometry Composer
.RPT
.PDF
.TPL
.CTF
.TXT
Acquisition
Date
File name
Sample ID
Ref date
Vol, mass
Uncertainty
Live time
Energy &
uncertainty
list
HV
Gain
Stabilizer
Filter
Library Editor
Nuclides
t½
keV
yield
uncertainty
Library
Nuclides
t½
keV
yield
uncertainty
Certificate Editor
PC clock
Cascade
summing
correction
File name
Dimensions
Materials
Densities
Detector SN
Template
Beaker
.BKR
.PAR
Detector
characteri-
zation
Dimensions
Materials
Beaker
Editor
QA Log
Report
Viewer
File name
Algorithm settings
Calibration
Report
template
Validation
parameters
Final report
Sample Container
.PDF
Data report
Test results

## Slide 8: Gammaspec 2016 sample
9,1698 (10) g
Fill-height ≈ 5 mm
Spectrum data file: GEC-987-K1-….cnf
Detector ID – Measurement ID – Acquisition cycle -

## Slide 9: Geometry
PLEX: Holder ID, absorber
GEC-987: Detector and measurement ID
WHI: Sample container ID
Unique name: GEC-987_PLEXWHI
Name is used for .GEO, .CAL,

## Slide 10: Geometry model (LabSocs)
Standard container ”WHITE.bkr”

## Slide 11
1. Assumed sample composition and weight entered
2. Density roughly adjusted to reach approx. fill-height

## Slide 12
Screenshots for documentation
Density automatically
calculated from
sample mass and fill height

## Slide 13
Genie report template (.tpl)
“miss”, * and @ are moved to the end
Section tag

## Slide 14
Nuclides listed with gaps!
Excel challenge…

## Slide 15: Genie algorithm settings

## Slide 16
QA
Documentation

## Slide 17
Genie algorithm settings QC
CAM-file parameter checksum

## Slide 18
Final Report

## Slide 19
A (w.m.) > MDA ?
Reporting rule
yes
no
“ A < [MDA]”
“ A (best est.) +/- u (best est.)”
To avoid false positive results, we have chosen not to report A when
MSA < A < MDA.
In some cases, w.m. < 0 and best. est. > 0.
w.m. = weighted mean
best. est. = best estimate (ISO11929)

## Slide 20: Loose end - Nuclide libraries
Data transfer?
Best data source?

## Slide 21
System validation
“Point”
Cs-137, Co-60, Am-241 (Ra-226, Pb-210)
“Beaker”
Eu-152
“Custom Beaker”
K-40 (KCl)

## Slide 22
QC parameter

## Slide 23
In case of non-compliance:
Increase efficiency uncertainty.
E.g. changed from factory default (4%)

## Slide 24
QC parameter
Energy tolerance (Eu-152)
Maximal deviation (keV)

## Slide 25: Drift
Energy sens.
Resolution
“Efficiency”

## Slide 26
Eu-152 LACE