## Slide 1: Sample measurement exercise (’PT’) Henrik Ramebäck Sofia Jonsson FOI Presentation@GAMMASPEC2016

## Slide 2: ’Measurement exercises’
What is your view of a measurement exercise, like a PT? (This is not an official PT though)
Is it something to avoid since there is a possibility that you ’fail’?
Is it an opportunity not only to show competence, but to pin-point possible problems in your measurement methods which then can be corrected?

## Slide 3: Task
Measure the activity of radionuclides in a sample!
Only five laboratories participated!!!
Six results reported for the activity of 137Cs [One lab reported two results: Different systems? Different locations?]
Two results reported for the activity of 40K Considered as background?

## Slide 4: The sample
IAEA-372, Radionuclides in grass:
40K (1060±56) Bq, k=2 137Cs (11320±360) Bq, k=2
Reference date: 2006-06-01
Activity 2016-09-09: 40K (1060±56) Bq, k=2 137Cs (8930±280) Bq, k=2

## Slide 5: A spectrum
Measured 73 ks using a low-background HPGe:

## Slide 6: Possible difficulties
’Pure’ empirical calibration:
-Biased efficiencies in close geometries
at ’high’ energies Possible problem for K-40!
-This results in biased effeciency also at lower energies: 7% at 662 keV in the
example Possible problem for Cs-137!
Unfortunately, only two participants measured 40K…
Biased efficiencies
due to TCS if measured in a
close geometry!

## Slide 7: Reporting
Six labs reported 137Cs
Two labs reported 40K
Four labs stated the coverage factor!
Two labs did not state the coverage factor (or confidence interval)

## Slide 8: Evaluation
Measurement results were evaluated using the ”ZETA-score”:
z<2: ”ok” [2<z<3: ”warning” z>3: ”not ok”]
For the two results not stating any coverage factor, the standard uncertainties were evaluated according to GUM assuming a rectangular distribution

## Slide 9: Results: 137Cs
z=7.5
z=2.1

## Slide 10: Results: 40K
Both labs reporting 40K were ok, i.e z<2!

## Slide 11
FOI CBRN Defence and Security, Umeå
?